## Slide 1: New approaches to sleep monitoring
Conor Heneghan PhD *,
Chief Engineer, ResMed Strategy and Ventures
Adjunct Associate Professor, School of Electrical, Electronic and Communication Engineering, University College Dublin
*Disclosure: Speaker is an employee and shareholder of ResMed

## Slide 2: Overview
Introduction
Current approaches to measuring sleep
Non-contact radio frequency biomotion sensing
Open research questions

## Slide 3: Who are ResMed?
Global company head quartered in San Diego, which provides therapy for sleep disordered breathing and other respiratory disorders primarily using non-invasive ventilation technology
Approximately 4000 people worldwide, $1.7B annual revenue
25-year old company, founded in Sydney, Australia
High emphasis on best-in-class products and the clinical research underpinning those products

## Slide 4: Sleep Disordered Breathing

## Slide 5: Sleep Disordered Breathing

## Slide 6: Why should we care about SDB?
Sleep-Disordered Breathing and Mortality: A Prospective Cohort Study, Punjabi et al., 2009, PLOS Medicine
Benefits of obstructive sleep apnoea treatment in coronary artery disease: a long-term follow-up study, Milleron et al., Eur Heart J., 2004

## Slide 7: Classical Sleep Measurement: Polysomnogram
Polysomnogram
Nasal airflow
Chest effort
Abdominal effort
Oximetry
EMG (legs and chin)
EOG
EEG (F4/M1, C4/M1, O2/M1)
ECG

## Slide 8: Classical Sleep Measurement: Polysomnogram
Waves | Frequency
Alpha | 7.5-12.5 Hz
Theta | 4-7 Hz
Delta | 0.5-4

## Slide 9: Classical Sleep Measurement: Polysomnogram
Gold standard sleep staging done by trained technician, reviewing
multiple EEG channels, and applying rules first formulated by Rechtshaffen and Kales in 1968 => AASM scoring manuals (2007)

## Slide 10: Classical Sleep Measurement: Polysomnogram
Berthomier at al., SLEEP, 2007.

## Slide 11: Single Channel EEG Measurements
96% agreement on sleep/wake scoring with expert consensus

## Slide 12: Consumer single channel EEG device

## Slide 13: Actigraphy approaches to sleep measurement

## Slide 14: Accuracy of Actigraphy

## Slide 15: Radio Frequency Biomotion Sensing

## Slide 16: Zone of Sensing: “Range Gated”

## Slide 17: Technical Characteristics
Characteristic | Value
Frequency | 10.525 GHz or 5.8 GHz
Emitted Power | ~1 mW
Clearances | FCC, CE, ARIB
Range | ~1.3m
3 dB Beamwidth | +/- 30°
Power consumption | ~60 mW

## Slide 18: Radio Frequency Biomotion Sensing

## Slide 19: Radio Frequency Biomotion Sensing

## Slide 20: Radio Frequency Biomotion Sensing

## Slide 21: Radio Frequency Biomotion Sensing

## Slide 22: Signal Processing
Motion Sensor
Movement / Motion
Separate into physiological components
Motion signal formation and digitisation
Core signal processing
Respiration
Attribute Extraction/ Pattern Recognition

## Slide 23: Detection of Movement

## Slide 24: Signal Processing Steps

## Slide 25: Initial Results on Sleep/Wake Staging

## Slide 26: Hypnograms
EEG
Non-contact sensor

## Slide 27: Latest Results on Sleep Wake Staging
Zaffaroni et al., European Sleep Research Society Annual Meeting, 2014

## Slide 28: ResMed S+ Product

## Slide 29: Radio Frequency Biomotion Sensing

## Slide 30: Respiratory Parameters: Detection of SDB

## Slide 31: Respiratory Parameters: Detection of SDB

## Slide 32: Respiratory Parameters: Respiration Rate

## Slide 33: Respiratory Parameters: Respiration Rate

## Slide 34: Open Questions
Ways to improve accuracy?
Sleep-onset latency estimation?
Validation across different population groups (severe insomniacs, heart failure, chronic back pain)
Validate if pharmacological effects on sleep EEG are seen in movement/respiration sleep measurements
What can we learn about population sleep characteristics?
Can this technology help in tracking patients with chronic diseases?

## Slide 35: Publications